scCO₂ Pressure Control System

Complete Implementation Guide

Taub Research Group

Department of Materials Science & Engineering

University of Michigan

March 2026

# Table of Contents

# 1. System Overview

This document provides a complete implementation guide for the supercritical CO₂ pressure control system. The system pressurizes a high-pressure vessel to 28 MPa at a controlled rate of 350 kPa/min, holds at the target pressure, and depressurizes at 2 MPa/min.

## 1.1 System Architecture

The system consists of three integrated layers:

- Simulink Digital Twin (Phase 1): A MATLAB/Simulink model that mirrors the physical plant for offline PID tuning before connecting to real hardware.

- Python Control Application (Phase 2): A real-time control loop running on Raspberry Pi with PID control, state machine, GUI, data logging, and cloud monitoring.

- ThingSpeak Cloud Dashboard (Phase 3): Remote monitoring via MathWorks ThingSpeak with live plots, email alerts, and MATLAB analytics.

## 1.2 Key Operating Parameters

| Parameter | Value | Notes |
| --- | --- | --- |
| Maximum Pressure | 28 MPa | Vessel rated limit |
| Pressurization Rate | 350 kPa/min | ~0.35 MPa/min via motorized valve |
| Depressurization Rate | 2 MPa/min | Via solenoid vent valve PWM |
| Hold Tolerance | ±0.3 MPa | Around target setpoint |
| PID Sample Time | 100 ms (10 Hz) | Control loop frequency |
| Temperature Pre-condition | ≥55°C | INKBIRD must reach setpoint before start |
| Emergency Rate Limit | 5 MPa/min | Immediate valve shutdown |
| Booster Pump Ratio | 62:1 | Air-driven, P_boost = 62 × P_drive |

## 1.3 State Machine

The control system operates through four states:

| State | Actuator | Control Action | Transition Condition |
| --- | --- | --- | --- |
| IDLE | All closed | Wait for operator START | Temperature ≥55°C + START button |
| PRESSURIZE | Motorized valve | PID tracks 350 kPa/min ramp setpoint | Pressure reaches target (28 MPa) |
| HOLD | Motorized valve | PID holds at target ±0.3 MPa | Operator presses STOP or timer expires |
| DEPRESSURIZE | Solenoid valve PWM | PID tracks 2 MPa/min down-ramp | Pressure < 0.5 MPa |

# 2. Simulink Digital Twin

The Simulink model replicates the physical plant dynamics so PID gains can be tuned offline before connecting to real hardware. This avoids dangerous trial-and-error tuning on a 28 MPa system.

## 2.1 Required Toolboxes

| Toolbox | Purpose | Required? |
| --- | --- | --- |
| MATLAB (base) | Script execution, MATLAB Function blocks | Yes |
| Simulink (base) | Block diagram simulation | Yes |
| Control System Toolbox | PID Tuner app, linearization | Yes |
| Simscape Fluids | Only if using physical domain blocks | Optional |

## 2.2 Model Block Diagram

The Simulink model (scCO₂_plant.slx) contains the following blocks connected in a closed-loop configuration:

Signal Flow

Ramp Setpoint ---->(+)----> PID Controller ----> Saturation [0,100%]

^  -                              |

|                                 v

Pressure Scope <--- | <---- Pressure Integrator <--- Pressure Plant

|                                 |

+--- measured pressure            |

v

Temperature Scope <-------- Temp Integrator <-------- Temp Plant

Block Details

| Block | Type | Configuration |
| --- | --- | --- |
| Clock | Source | Simulation time input |
| Ramp Rate Gain | Gain | 0.35/60 = 0.005833 MPa/s |
| Sum (setpoint) | Sum | Adds ramp to initial pressure |
| MinMax (cap at 28) | MinMax | Clamps setpoint at 28 MPa |
| PID Controller | Discrete PID | Ts=0.1s, anti-windup clamping |
| Saturation | Saturation | Limits valve output to [0, 100]% |
| Pressure Plant | MATLAB Function | dP = f(valve, P_booster, P_vessel) |
| Temp Plant | MATLAB Function | dT = compression + J-T + INKBIRD |
| Pressure Integrator | Discrete Integrator | IC=0.1 MPa (ambient) |
| Temp Integrator | Discrete Integrator | IC=25°C (room temp) |
| White Noise | Random Source | Sensor noise: σ=0.005 MPa |
| Pressure Scope | Scope | Shows pressure + setpoint |
| Temperature Scope | Scope | Shows temperature over time |
| Valve Scope | Scope | Shows valve % output |

## 2.3 Plant Equations

The MATLAB Function blocks implement these physics equations:

Pressure Plant

function dP = pressure_plant(valve_pct, P_vessel, P_booster)

K_FLOW = 0.003;

delta_P = max(P_booster - P_vessel, 0);

dP = K_FLOW * valve_pct * sqrt(delta_P);

end

Temperature Plant

function dT = temp_plant(dP_flow, dP_vent, T, P_booster)

K_COMPRESS = 0.8;  K_JT = 0.25;  K_INKBIRD = 0.067;

T_setpoint = 55;  % INKBIRD target

dT_compress = K_COMPRESS * max(dP_flow, 0) / P_booster;

dT_jt       = K_JT * min(dP_vent, 0);

dT_inkbird  = -K_INKBIRD * (T - T_setpoint);

dT = dT_compress + dT_jt + dT_inkbird;

end

## 2.4 Build Procedure

Two options are available to build the Simulink model:

Option A: Automated Build (Recommended)

- Open MATLAB and navigate to the simulink/ folder.

- Run build_model.m in the MATLAB Command Window. This script creates scCO₂_plant.slx with all blocks, connections, and initial PID gains (Kp=40, Ki=5, Kd=2).

- Run run_simulation.m to execute a 5000-second simulation and generate plots.

- Verify the pressure ramp rate is approximately 350 kPa/min in the console output.

Option B: Manual Build

If you prefer to build the model by hand, follow these steps:

- Create a new Simulink model: File > New > Blank Model. Save as scCO₂_plant.slx.

- Add a Clock block (Sources library) and connect it to a Gain block set to 0.35/60 (= 0.005833 MPa/s).

- Add a Sum block, then a MinMax block (set to min, upper limit = 28). This creates the ramp setpoint clamped at 28 MPa.

- Add a Discrete PID Controller block. Set sample time Ts=0.1, P=40, I=5, D=2, filter coefficient N=100. Enable anti-windup (clamping). Set output saturation [0, 100].

- Add a Saturation block with limits [0, 100] after the PID output.

- Add two MATLAB Function blocks for Pressure Plant and Temperature Plant using the equations above.

- Add two Discrete-Time Integrator blocks: Pressure (IC=0.1 MPa) and Temperature (IC=25°C). Set sample time to 0.1 for both.

- Add a Band-Limited White Noise block (noise power = 0.005², sample time = 0.1) and sum it with the pressure output for sensor noise simulation.

- Connect all blocks as shown in the signal flow diagram above. Add Scope blocks for pressure, temperature, and valve output.

## 2.5 PID Tuning Procedure

After building and running the initial simulation:

- Double-click the PID Controller block in Simulink.

- Click the "Tune..." button to open the PID Tuner app.

- The app will linearize the plant around the current operating point and display a Bode plot.

- Use the Response Time slider (move right for faster response, left for less overshoot).

- Use the Transient Behavior slider to trade off between aggressive and robust tuning.

- Target specifications: Rise time < 120s, Overshoot < 5%, Settling time < 300s.

- Click "Update Block" to write the tuned gains back to the PID Controller block.

- Re-run the simulation to verify the pressure ramp rate is approximately 350 kPa/min.

- Run export_gains.m to save the tuned Kp, Ki, Kd values to config_gains.mat and get a copy-paste block for Python config.py.

Tuning Targets

| Metric | Target | Acceptable Range |
| --- | --- | --- |
| Rise time (0→28 MPa) | < 120 seconds | 90–150 seconds |
| Overshoot | < 5% | 0–8% |
| Steady-state error | < 0.1 MPa | < 0.3 MPa |
| Ramp rate (pressurize) | 350 kPa/min | 300–400 kPa/min |
| Settling time | < 300 seconds | < 500 seconds |

# 3. Python Control Application

The Python application is the real-time control system that runs on the Raspberry Pi (or in simulation mode on any Windows/Linux/Mac machine).

## 3.1 Software Prerequisites

| Software | Version | Purpose |
| --- | --- | --- |
| Python | 3.9+ | Runtime environment |
| pip packages | matplotlib, requests | GUI plotting, ThingSpeak HTTP |
| tkinter | Built-in with Python | GUI framework |
| RPi.GPIO | Latest (RPi only) | GPIO pin control |
| adafruit-circuitpython-ads1x15 | Latest (RPi only) | ADS1115 ADC for pressure sensor |

## 3.2 Installation

# Clone or copy the pressure_control/ folder to the target machine

cd pressure_control/

# Install dependencies

pip install -r requirements.txt

# Run in simulation mode (no hardware needed)

python main.py

## 3.3 Configuration

All tunable parameters are in config.py. Key settings:

| Parameter | Default | Description |
| --- | --- | --- |
| SIMULATION | True | True = simulated physics, False = real RPi GPIO |
| TARGET_PRESSURE_MPA | 28.0 | Final target pressure |
| PRESSURIZE_RATE_KPA_MIN | 350.0 | Ramp-up rate |
| DEPRESSURIZE_RATE_MPA_MIN | 2.0 | Ramp-down rate |
| PRESS_KP / KI / KD | 40.0 / 5.0 / 2.0 | PID gains for pressurization |
| HOLD_KP / KI / KD | 30.0 / 8.0 / 1.0 | PID gains for hold state |
| DEPRESS_KP / KI / KD | 20.0 / 3.0 / 1.0 | PID gains for depressurization |
| SAMPLE_TIME_S | 0.1 | PID update interval (100 ms) |
| TEMP_MIN_START_C | 55.0 | Minimum temperature to allow START |
| THINGSPEAK_API_KEY | (placeholder) | Write API key for cloud upload |

## 3.4 Application Architecture

The application uses a multi-threaded architecture:

- Main Thread: Runs the Tkinter GUI event loop. Refreshes display every 200 ms via root.after() callback.

- Control Thread: Runs the PID control loop at 10 Hz. Reads sensors, computes PID output, writes actuators, logs data.

- Simulator Thread: (Simulation mode only) Runs physics model at 10 Hz in background, integrating pressure and temperature.

- Cloud Thread: Uploads data to ThingSpeak every 2 seconds in background. Gracefully disabled if API key is not set.

## 3.5 Hardware Abstraction Layer (HAL)

The HAL pattern allows the same control code to run on both simulated and real hardware:

# In config.py, set SIMULATION = True for testing, False for RPi

# hardware/base.py    - Abstract base class (interface)

# hardware/simulator.py - Simulated physics (Windows/Mac/Linux)

# hardware/rpi_hardware.py - Real GPIO + ADS1115 (Raspberry Pi)

# Swap is automatic based on SIMULATION flag in main.py:

#   if SIMULATION:

#       hw = SimulatedHardware()

#   else:

#       hw = RPiHardware()

## 3.6 Safety Features

- Rate limiter with 5-second rolling window averages pressure rate of change.

- Warning logged at 600 kPa/min; emergency valve shutdown at 5 MPa/min.

- Minimum pressure threshold (1 MPa) to ignore rate noise at ambient.

- E-stop button on GUI triggers immediate valve closure in any state.

- Temperature pre-condition: START button disabled if temperature < 55°C.

- PID output clamped to [0, 100]% with anti-windup on integral term.

- Derivative computed on measurement (not error) to avoid derivative kick on setpoint changes.

# 4. ThingSpeak Cloud Monitoring

ThingSpeak provides free cloud data logging (3 million messages/year) with built-in visualization, MATLAB analytics, and email alerts.

## 4.1 Account Setup

- Go to thingspeak.com and create a free MathWorks account.

- Click "New Channel" and name it "scCO2 Pressure Control".

- Configure 4 fields: Field 1 = Pressure (MPa), Field 2 = Temperature (°C), Field 3 = Setpoint (MPa), Field 4 = State Code.

- Go to API Keys tab and copy the Write API Key.

- Paste the key into config.py as THINGSPEAK_API_KEY.

## 4.2 State Codes

| Code | State | Description |
| --- | --- | --- |
| 0 | IDLE | System idle, waiting for operator |
| 1 | PRESSURIZE | Ramping up at 350 kPa/min |
| 2 | HOLD | Maintaining target pressure |
| 3 | DEPRESSURIZE | Ramping down at 2 MPa/min |
| 99 | E-STOP | Emergency stop triggered |

## 4.3 Email Alerts (React App)

ThingSpeak React apps can send email alerts on abnormal conditions:

- In your ThingSpeak channel, go to Apps > React.

- Create a new React: Condition = "Field 1 (Pressure) > 29" (over-pressure).

- Action = "ThingHTTP" pointing to an email service, or use MATLAB Analysis to send email.

- Optionally add a second React for State Code = 99 (E-stop triggered).

## 4.4 MATLAB Integration

Read data from ThingSpeak directly in MATLAB for post-experiment analysis:

% Read last 1000 data points from channel

channelID = 123456;  % Replace with your channel ID

readAPIKey = 'YOUR_READ_KEY';

data = thingSpeakRead(channelID, 'NumPoints', 1000, ...

'ReadKey', readAPIKey);

plot(data.Timestamps, data.Pressure);

xlabel('Time'); ylabel('Pressure (MPa)');

# 5. Hardware Bill of Materials

The following components are required to build the physical pressure control system:

## 5.1 Electronics

| Item | Quantity | Specification | Purpose |
| --- | --- | --- | --- |
| Raspberry Pi 4B | 1 | 4 GB RAM, 32 GB SD card | Main controller |
| ADS1115 ADC | 1 | 16-bit, 4-channel, I²C | Analog pressure reading |
| Pressure Transducer | 1 | 0–28 MPa, 4–20 mA output | Vessel pressure measurement |
| 250Ω Shunt Resistor | 1 | 0.1% precision | 4–20 mA to 1–5V conversion |
| INKBIRD Temperature Controller | 1 | Model ITC-106VH or similar | Vessel temperature control |
| Motorized Ball Valve | 1 | 0–5V or 4–20 mA control | Pressurization flow control |
| Solenoid Valve | 1 | 24 VDC, normally closed | Depressurization vent |
| MOSFET Module (IRF520) | 1 | N-channel, logic level | Solenoid PWM driver |
| 24V DC Power Supply | 1 | 2A minimum | Solenoid valve power |
| 5V DC Power Supply | 1 | 3A (RPi + peripherals) | Raspberry Pi power |
| E-Stop Button | 1 | Normally closed, panel mount | Emergency shutdown |
| DIN Rail Terminal Blocks | 8 | Various | Wiring junction points |

## 5.2 Pneumatic / Hydraulic

| Item | Quantity | Specification | Purpose |
| --- | --- | --- | --- |
| Air-Driven Booster Pump | 1 | 62:1 ratio, rated 28 MPa | Pressure source |
| High-Pressure Vessel | 1 | Rated > 28 MPa, 316 SS | Process chamber |
| HP Tubing (1/4" OD) | As needed | 316 SS, rated 40 MPa | Pressure lines |
| HP Fittings (Swagelok) | As needed | 316 SS, 1/4" tube | Connections |
| Check Valve | 1 | Rated > 28 MPa | Prevent backflow |
| Relief Valve | 1 | Set at 30 MPa | Over-pressure safety |

# 6. Wiring Guide

## 6.1 Raspberry Pi GPIO Pin Assignment

| GPIO Pin | Direction | Connection | Notes |
| --- | --- | --- | --- |
| GPIO 18 (PWM0) | Output | Motorized valve signal | Hardware PWM for smooth control |
| GPIO 13 (PWM1) | Output | MOSFET gate (solenoid) | PWM duty cycle 0–100% |
| GPIO 17 | Input | E-stop button | Pull-up enabled, active LOW |
| GPIO 2 (SDA) | I²C | ADS1115 SDA | I²C data line |
| GPIO 3 (SCL) | I²C | ADS1115 SCL | I²C clock line |
| 3.3V | Power | ADS1115 VDD | ADC power supply |
| GND | Ground | Common ground | All grounds tied together |

## 6.2 ADS1115 ADC Connections

| ADS1115 Pin | Connection | Notes |
| --- | --- | --- |
| VDD | RPi 3.3V | Power supply |
| GND | RPi GND | Common ground |
| SDA | RPi GPIO 2 | I²C data |
| SCL | RPi GPIO 3 | I²C clock |
| A0 | Shunt resistor (+) | Pressure signal (1–5V) |
| ADDR | GND | I²C address = 0x48 |

## 6.3 Pressure Sensor Circuit

The pressure transducer outputs 4–20 mA. A 250Ω shunt resistor converts this to 1–5V for the ADS1115:

Pressure Transducer (4-20 mA)

|

+-----[ 250 ohm ]-----+--- ADS1115 A0 (+)

|                      |

GND                    GND--- ADS1115 GND

Voltage = Current x 250 ohm

At  0 MPa:  4 mA x 250 = 1.0 V

At 28 MPa: 20 mA x 250 = 5.0 V

Conversion: P(MPa) = (V - 1.0) / (5.0 - 1.0) x 28.0

## 6.4 Solenoid Valve Driver Circuit

RPi GPIO 13 ----> MOSFET Gate (IRF520 module)

24V Supply  ----> MOSFET Drain ----> Solenoid (+)

Solenoid (-) ----> MOSFET Source ----> GND

Note: IRF520 module includes flyback diode.

If using bare MOSFET, add 1N4007 diode across solenoid.

# 7. Pressure Tubing & Plumbing

## 7.1 System Schematic

Shop Air (0.6 MPa)

|

v

[Air-Driven Booster Pump 62:1]

|

| (High-pressure output ~37 MPa max)

v

[Check Valve] ---> [Motorized Ball Valve] ---> [Vessel]

|

[Relief Valve 30 MPa] <---+

|

[Solenoid Vent Valve] <---+

|

[Pressure Transducer] <---+

|

[INKBIRD Thermocouple] <--+

## 7.2 Tubing Requirements

- All high-pressure lines: 1/4" OD x 0.065" wall 316 stainless steel tubing.

- Rated working pressure: 40 MPa minimum (provides 1.4x safety factor over 28 MPa operating).

- Fittings: Swagelok or equivalent compression fittings, 316 SS.

- Vent line from solenoid valve: route to fume hood or exterior exhaust.

- Minimize tubing length to reduce dead volume and improve response time.

## 7.3 Assembly Notes

- Cut tubing with a tube cutter (not a hacksaw) for clean, burr-free ends.

- Deburr all cut ends with a deburring tool.

- Assemble Swagelok fittings per manufacturer torque specifications (1-1/4 turns past finger-tight for new fittings).

- Pressure test the complete system at 1.25x operating pressure (35 MPa) with inert gas before first use.

- Check all fittings with leak detection fluid (Snoop or soapy water) at test pressure.

- Install the relief valve BEFORE the motorized valve so it protects the vessel even if the valve fails open.

# 8. Raspberry Pi Deployment

## 8.1 OS Setup

- Flash Raspberry Pi OS (64-bit, Bookworm) to a 32 GB microSD card using Raspberry Pi Imager.

- Enable SSH and set up Wi-Fi during imaging (Advanced Options).

- Boot the RPi, connect via SSH, and run: sudo apt update && sudo apt upgrade

- Enable I²C: sudo raspi-config > Interface Options > I2C > Enable

- Install Python dependencies: pip install matplotlib requests adafruit-circuitpython-ads1x15 RPi.GPIO

## 8.2 Application Deployment

- Copy the pressure_control/ folder to the RPi (via scp, USB, or git clone).

- Edit config.py: set SIMULATION = False.

- Set THINGSPEAK_API_KEY to your Write API key.

- Verify I²C connection: i2cdetect -y 1 (should show device at 0x48).

- Test: python main.py (GUI requires display; use VNC or direct HDMI monitor).

## 8.3 Auto-Start on Boot (Optional)

To start the application automatically when the RPi boots:

# Create a systemd service file

sudo nano /etc/systemd/system/scco2.service

[Unit]

Description=scCO2 Pressure Control

After=network.target

[Service]

User=pi

WorkingDirectory=/home/pi/pressure_control

ExecStart=/usr/bin/python3 main.py

Restart=on-failure

Environment=DISPLAY=:0

[Install]

WantedBy=multi-user.target

# Enable and start the service

sudo systemctl enable scco2.service

sudo systemctl start scco2.service

# 9. Commissioning Procedure

Follow these phases in order when bringing the system online for the first time.

## 9.1 Phase 1: Electrical Verification (No Pressure)

- Verify all wiring against the GPIO pin assignment table (Section 6.1).

- Power on RPi. Run: python -c "import board; import busio; i2c = busio.I2C(board.SCL, board.SDA); print(i2c.scan())" to confirm ADS1115 at address 72 (0x48).

- Test motorized valve: manually command 0%, 50%, 100% via GUI and verify physical valve position.

- Test solenoid valve: command PWM duty and listen for click / verify flow with low-pressure air.

- Test E-stop button: press button and verify GUI shows E-STOP state and all valves close.

- Verify pressure transducer reads ambient (~0.1 MPa) with no pressure applied.

## 9.2 Phase 2: Low-Pressure Test (< 5 MPa)

- Connect booster pump with shop air at reduced pressure (e.g., 0.05 MPa drive = ~3 MPa boost).

- Start the application. Set target pressure to 3 MPa in config.py temporarily.

- Press START and observe pressurization. Verify rate is approximately 350 kPa/min.

- Verify HOLD state maintains pressure within tolerance.

- Press STOP and verify controlled depressurization at 2 MPa/min.

- Check ThingSpeak channel for uploaded data points.

- Review CSV log file in logs/ folder for data integrity.

## 9.3 Phase 3: Full-Pressure Commissioning

- Restore target pressure to 28 MPa in config.py.

- Ensure relief valve is installed and set to 30 MPa.

- Ensure INKBIRD controller is active and heating vessel to 55°C.

- Increase drive air pressure to full shop air (0.6 MPa).

- Run a complete pressurize-hold-depressurize cycle while monitoring all readings.

- If PID tuning is poor (oscillation, overshoot > 8%, slow response), adjust gains in config.py. Use the Simulink model to test new gains before applying.

## 9.4 Phase 4: Validation

- Run 3 consecutive full cycles without manual intervention.

- Verify pressurization rate stays within 300–400 kPa/min for each cycle.

- Verify hold pressure stays within ±0.3 MPa of target.

- Verify depressurization rate stays within 1.5–2.5 MPa/min.

- Verify no E-stop triggers or rate limit warnings during normal operation.

- Download ThingSpeak data and compare with local CSV logs for consistency.

# 10. Troubleshooting

| Symptom | Likely Cause | Solution |
| --- | --- | --- |
| Pressure not rising | Motorized valve not opening | Check GPIO 18 output; verify valve responds to 0–5V signal |
| Pressure rises too slowly | PID gains too low or booster supply pressure low | Increase Kp; verify shop air is at full 0.6 MPa |
| Pressure overshoots target | PID gains too aggressive (high Kp) | Reduce Kp by 20–30%; increase Kd; re-tune in Simulink |
| Pressure oscillates in HOLD | Integral windup or gains too high | Reduce Ki; verify anti-windup is active; check for leaks |
| Rate limit warning at startup | Noise at low pressure (< 1 MPa) | Normal — warning suppressed below 1 MPa threshold |
| E-stop triggers unexpectedly | Wiring issue or noise on GPIO 17 | Check E-stop wiring; add 10k pull-up resistor; verify debounce |
| ADS1115 not detected | I²C not enabled or wiring error | Run sudo raspi-config to enable I²C; check SDA/SCL connections |
| ThingSpeak not receiving data | API key not set or network issue | Verify THINGSPEAK_API_KEY in config.py; check Wi-Fi connection |
| GUI not launching | Display not available | Use VNC or connect HDMI monitor; set DISPLAY=:0 in environment |
| CSV log file empty | logs/ directory missing | Create logs/ folder in pressure_control/; check write permissions |

# 11. File Structure Reference

| File | Purpose |
| --- | --- |
| pressure_control/main.py | Application entry point; builds hardware, state machine, GUI; starts control thread |
| pressure_control/config.py | All tunable constants: PID gains, physics parameters, GPIO pins, ThingSpeak key |
| pressure_control/hardware/base.py | Abstract hardware interface (HAL base class) |
| pressure_control/hardware/simulator.py | Simulated physics engine with background thread (Windows/Mac/Linux) |
| pressure_control/hardware/rpi_hardware.py | Real hardware interface: GPIO outputs, ADS1115 pressure input |
| pressure_control/control/pid.py | PID controller with derivative-on-measurement and anti-windup |
| pressure_control/control/ramp.py | Linear ramp setpoint generator (ascending/descending) |
| pressure_control/control/state_machine.py | IDLE/PRESSURIZE/HOLD/DEPRESSURIZE state machine with rate limiter |
| pressure_control/logger.py | CSV data logger (one file per experiment run) |
| pressure_control/cloud.py | ThingSpeak HTTP uploader (background thread, 2-second interval) |
| pressure_control/gui.py | Tkinter + Matplotlib GUI with live pressure/temperature plots |
| simulink/build_model.m | Creates scCO₂_plant.slx Simulink model programmatically |
| simulink/run_simulation.m | Runs 5000-second simulation and generates plots |
| simulink/export_gains.m | Exports tuned PID gains from Simulink to .mat and prints config.py block |

End of Implementation Guide.