## Slide 1
Automated scCO₂
Pressure Regulation System
Closed-loop computer control for supercritical CO₂ research
Raspberry Pi · Python PID · Peng-Robinson EOS
Mechanical Engineering Research | University of Michigan–Dearborn | 2025

## Slide 2
What is Supercritical CO₂?
CO₂ becomes supercritical above 31.1°C and 7.38 MPa
Acts like a liquid and a gas at the same time — high density, low viscosity
Dissolves materials like a liquid but flows like a gas
Used in extraction, sterilization, and materials processing
Non-toxic, non-flammable, and leaves no residue
SOLID
LIQUID
GAS
SUPERCRITICAL
FLUID ★
← Pressure
Tc = 31.1°C Pc = 7.38 MPa
Our system operates at 60°C, up to 28 MPa — deep in the supercritical zone

## Slide 3
System Performance Targets
28 MPa
Max Operating
Pressure
4,061 PSI — vessel
rated to 34.5 MPa
350 kPa/min
Pressurization
Rate
Controlled ramp
to avoid thermal shock
±0.30 MPa
Hold Accuracy
Band
Maintained by
dedicated hold PID
60°C
Operating
Temperature
Above Tc to ensure
supercritical phase
All controlled autonomously — operator sets target, system handles the rest

## Slide 4
System Architecture
CO₂
Cylinder
Bone Dry
5.7 MPa
▶
Air-Driven
Booster
HII 5G-TD
172 MPa max
▶
Needle
Valve
Stepper motor
controlled
▶
PARR
Vessel
1L, 316SS
34.5 MPa rated
▶
Vent
Solenoid
PWM
controlled
Raspberry Pi 4
Python PID Control | 10 Hz | GUI + Logging
Pressure
Sensor
Yellow = sensor feedback | Teal = control signal | White arrows = gas flow

## Slide 5
Parts List
Component | Specification | Purpose
PARR Pressure Vessel | 1L, 316SS, 5000 PSI, Model 2302HC | Contains experiment
HII Gas Booster | 5G-TD-28/150-CO2, 25,000 PSI max out | Amplifies CO₂ to operating pressure
CO₂ Cylinder | Bone Dry, UN1013, 50 lb | Process gas supply (≥99.8% pure)
Motorized Needle Valve | Stepper-driven, high pressure rated | Controls pressurization flow rate
Vent Solenoid Valve | PWM-controlled, normally closed | Controls depressurization rate
Safety Relief Valve | Parker SS316, 1/4" MNPT | Hardware overpressure protection
Pressure Transducer | Ashcroft 0–5000 PSI, 4–20 mA, IP67 | Real-time pressure feedback
INKBIRD Controller | Setpoint 60°C, stable ±0.01°C | Vessel temperature control
Raspberry Pi 4 | Python 3.11, 10 Hz control loop | Main controller
ADS1115 ADC | 16-bit, I²C, 4-channel | Reads 4–20 mA transducer signal
Stepper Motor + Driver | NEMA 17, A4988, 1/16 microstepping | Actuates needle valve
DIN Rail PSU | Dayton 33NT20, 24VDC 50W | 24V system power
Relay Module | 24VDC coil, 10A contacts | GPIO drives solenoid valve

## Slide 6
Automated Control Logic
IDLE
Valves closed
Awaiting start
Temp check ≥55°C
▶
PRESSURIZE
PID ramps pressure
350 kPa/min
Kp=40 Ki=5 Kd=2
▶
HOLD
PID holds setpoint
±0.30 MPa band
Kp=20 Ki=2
▶
DEPRESSURIZE
PID vents slowly
2 MPa/min
Kp=60 Ki=8
Safety Interlocks (active in ALL states)
Overpressure cutoff at 28.5 MPa — forces IDLE immediately
Rate limiter: closes all valves if pressure rises > 5 MPa/min
Temperature gate: pressurization blocked below 55°C
Watchdog: forces IDLE if control loop stalls for > 1 second

## Slide 7
Physics Model — Real CO₂ Behavior
Peng-Robinson Equation of State
At 20+ MPa, CO₂ is up to 118% denser than ideal gas predicts
The simulator uses PR-EOS so pressure predictions are physically accurate
A 0.01°C temperature flicker moves pressure by 3.5 kPa at 20 MPa — not 0.6 kPa as ideal gas predicts
This matters most in the HOLD phase where stability is critical
Z-factor & Density at 60°C
Pressure | Z-factor | Density vs Ideal
0.34 MPa | 0.99 | ≈ same
5 MPa | 0.79 | +26% denser
10 MPa | 0.54 | +84% denser
20 MPa | 0.46 | +118% denser ★
28 MPa | 0.55 | +82% denser
★ At 20 MPa, ideal gas is off by 2× — PR-EOS corrects this

## Slide 8
Software Stack
main.py
Entry point, control thread, watchdog
control/pid.py
PID with anti-windup, derivative on measurement
control/ramp.py
Linear ramp setpoint generator
control/eos.py
Peng-Robinson EOS — real CO₂ physics
control/state_machine.py
IDLE/PRESSURIZE/HOLD/DEPRESSURIZE logic
hardware/simulator.py
Windows physics simulation (10 Hz)
hardware/rpi_hardware.py
Raspberry Pi GPIO/ADS1115 (hardware stub)
gui.py
Tkinter + Matplotlib live dashboard
logger.py
CSV timestamped experiment logs
cloud.py
ThingSpeak real-time cloud monitoring

## Slide 9
Next Steps
01
Complete RPi Hardware Driver
Wire ADS1115 for pressure transducer, stepper GPIO for needle valve, relay for solenoid — the final step before real hardware runs.
In Progress
02
PID Gain Re-Tuning
PR-EOS shows CO₂ is 5-6× more pressure-sensitive to temperature at 20+ MPa. Hold PID gains need verification in updated simulation.
Pending
03
Valve Cv Calibration
Run timed pressurization at fixed valve opening, back-calculate Cv from pressure rise + PR-EOS density to remove last empirical constant.
Pending
04
Simulink Digital Twin
Use build_model.m + PID Tuner with Simscape Fluids to get analytically-tuned gains. Export to config.py via export_gains.m.
Planned

## Slide 10
Summary
Designed and built a fully automated scCO₂ pressure control system from scratch
Implemented 3-PID state machine (PRESSURIZE / HOLD / DEPRESSURIZE) running at 10 Hz on Raspberry Pi
Physics simulator uses Peng-Robinson EOS — captures real CO₂ behavior at 20–28 MPa with 5–6× better accuracy than ideal gas
Hardware abstraction layer allows the same Python code to run in Windows simulation or on real RPi hardware
Real-time GUI dashboard + CSV logging + ThingSpeak cloud monitoring
System validated in simulation against PARR vessel (1L, 34.5 MPa rated) and HII booster (172 MPa max output)
Questions?