## Slide 1
SUPERCRITICAL CO₂ PRESSURE CONTROL SYSTEM
Automated Pressurization & Depressurization Rate Control
University of Michigan Materials Science & Engineering Taub Research Group
March 2026

## Slide 2
THE CHALLENGE
Current System: Manual pressure control with no rate regulation
No Constant Rate
Pressurization rate varies unpredictably based on manual valve adjustment
Operator Dependent
Results vary based on who runs the experiment and how attentive they are
No Data Logging
Pressure vs time data must be recorded manually or not at all
Safety Risk
No automated overpressure protection or controlled depressurization
Research Requirement: Constant, reproducible pressurization rate of 350 kPa/min

## Slide 3
SYSTEM REQUIREMENTS
Parameter
Specification
Notes
Maximum Pressure
28 MPa (4000 psi)
Vessel design limit
Operating Temperature
60°C (140°F)
Above CO₂ critical point
Pressurization Rate
350 kPa/min
Constant, controllable
Depressurization Rate
2 MPa/min
Constant, controllable
Rate Accuracy
±5-10%
Acceptable for research
Hold Stability
±1-2%
During treatment phase
Control Interface
Remote/Automated
Start, stop, set parameters
Key Insight: At 60°C, CO₂ is supercritical — single phase, no liquid-vapor transitions

## Slide 4
OPTION A: COMMERCIAL SOLUTION
Emerson TESCOM ER5000 + Dome-Loaded Regulator
ADVANTAGES
• Turnkey solution — minimal engineering • Industrial-grade reliability • Proven for high-pressure gas service • Technical support from Emerson
DISADVANTAGES
• Cost: $8,000 - $15,000+ • Lead time: 6-8 weeks for configured valves • Proprietary software (ERtune) • Limited data logging capability • No integrated thermodynamics for scCO₂
Total Estimated Cost: $12,000 - $15,000 | Timeline: 3-4 months

## Slide 5
OPTION B: DIY CONTROL SYSTEM
Custom-built solution using industrial components + microcontroller/DAQ
ADVANTAGES
• Cost: $2,500 - $3,500 (80% savings) • Timeline: 4-6 weeks • Full control over software and data • Integrated data logging for research • Expandable and customizable • Learning opportunity for lab
CONSIDERATIONS
• Requires integration effort • Rate accuracy: ±5-10% (vs ±2% commercial) • Documentation and maintenance on us • Some trial-and-error in tuning
Total Estimated Cost: $2,500 - $3,500 | Timeline: 4-6 weeks

## Slide 6
PROPOSED SYSTEM ARCHITECTURE
Direct CO₂ flow control between booster and vessel
CO₂ Cylinder
Victor Regulator
Air-Driven Booster
MOTORIZED NEEDLE VALVE
Pressure Vessel
→
→
→
→
CONTROL SYSTEM
• Pressure Transducer → reads vessel pressure (4-20mA) • Controller (RPi/LabVIEW) → runs PID algorithm • Stepper Motor → adjusts needle valve position • PWM Solenoid → controls depressurization
↑ Feedback
KEY: Motorized valve provides DIRECT control of CO₂ flow to vessel

## Slide 7
P&ID SCHEMATIC
PRESSURIZATION PATH:
CO₂ Tank (850 psi) → Victor Regulator (800 psi out) → Air-Driven Booster (28 MPa out) → MOTORIZED NEEDLE VALVE → Check Valve → Pressure Vessel
DEPRESSURIZATION PATH:
Pressure Vessel → PWM Solenoid Valve → Manual Needle Valve (flow limiter) → Vent to Fume Hood
CONTROL LOOP:
Pressure Transducer (4-20mA) → Signal Conditioner → Controller (PID) → Stepper Driver → Motorized Valve ↘ PWM Signal → SSR → Vent Solenoid
COMPONENT LOCATIONS
┌─────────────────────────────────────┐ │ │ │ [CO₂ TANK] → [VICTOR REG] │ │ ↓ │ │ [BOOSTER] → [MOTORIZED VALVE] │ │ ↓ ↓ │ │ [AIR REG] [PRESSURE VESSEL] │ │ ↓ │ │ [PWM SOLENOID] │ │ ↓ │ │ [VENT] │ │ │ └─────────────────────────────────────┘
Note: Existing equipment in GRAY, new components in GREEN

## Slide 8
HOW THE CONTROL LOOP WORKS
PRESSURIZE
Controller opens motorized valve to allow CO₂ flow. PID adjusts valve position to maintain dP/dt = 350 kPa/min until target.
HOLD
Target pressure reached. Controller maintains setpoint by small adjustments. Vessel held at 28 MPa, 60°C for treatment.
DEPRESSURIZE
PWM solenoid opens in pulses to vent CO₂ at controlled rate. PID adjusts duty cycle for -dP/dt = 2 MPa/min.
PID Control Algorithm
Output = Kp × Error + Ki × ∫Error + Kd × dError/dt
Error = Setpoint - Measured Pressure | Setpoint ramps linearly: P(t) = P₀ + (rate × time) | Controller adjusts valve until Error → 0
Pressurize 0→28 MPa: ~80 min | Hold: variable | Depressurize 28→0 MPa: ~14 min

## Slide 9
BILL OF MATERIALS
Component
Part Number / Description
Est. Cost
Motorized Needle Valve
HiP valve + NEMA 17 stepper + coupling
$800-1,200
HP Solenoid (depress)
Clark Cooper EH40-04-D024-XP
$600-700
Pressure Transducer
Omega PX309-5KGI (0-35 MPa)
$300-350
Needle Valve (vent)
Swagelok SS-1RS4-A
$150-180
Check Valve
Swagelok SS-4C-1/3
$80-100
HP Fittings & Tubing
Various Swagelok 1/4" HP
$250-300
Controller (Option A)
Raspberry Pi 4 + ADC/DAC HAT
$150-200
Controller (Option B)
NI USB-6001 + LabVIEW
$400-500
Stepper Driver & PSU
TB6600 + Mean Well 24V
$80-100
Wiring, Enclosure, Misc
Various electrical
$100-150
TOTAL: $2,500 - $3,500

## Slide 10
CONTROLLER OPTIONS
RASPBERRY PI + PYTHON
Hardware: • Raspberry Pi 4 (4GB RAM) • ADS1115 ADC (4-20mA input) • MCP4725 DAC (analog output) • TB6600 Stepper Driver Software: • Python with simple-pid library • Custom GUI (Tkinter/PyQt) • CSV data logging built-in Cost: ~$200 | Flexible & customizable
NI DAQ + LABVIEW
Hardware: • NI USB-6001 or CompactDAQ • Built-in AI/AO/DIO channels • Industrial-grade signal conditioning Software: • LabVIEW (UMich has licenses) • Built-in PID toolkit • Professional GUI tools • TDMS data logging Cost: ~$500 | Industry standard
Recommendation: Start with Raspberry Pi for flexibility; upgrade to LabVIEW if needed

## Slide 11
RASPBERRY PI IMPLEMENTATION
WIRING CONNECTIONS
PRESSURE TRANSDUCER (4-20mA): Wire+ → 250Ω resistor → ADS1115 A0 Wire- → GND (common with RPi) Result: 1-5V input to ADC STEPPER MOTOR (TB6600 driver): TB6600 PUL+ → GPIO 17 TB6600 DIR+ → GPIO 27 TB6600 ENA+ → GPIO 22 All negative → GND Motor → TB6600 A+/A-/B+/B- PWM SOLENOID: GPIO 18 → SSR input (+) GND → SSR input (-) SSR output → 24V solenoid E-STOP: NC button → GPIO 4 (w/ pullup) Also wired to power relay
SOFTWARE MODULES
main.py ├── config.py # Constants, pins, PID params ├── sensors.py # ADC reading, calibration ├── actuators.py # Stepper control, PWM ├── pid.py # PID controller logic ├── state_machine.py # IDLE/PRESS/HOLD/DEPRESS ├── safety.py # E-stop, limits, watchdog ├── logger.py # CSV data recording └── gui.py # Tkinter interface Key Libraries: • simple-pid (PID control) • adafruit-ads1x15 (ADC) • RPi.GPIO (IO control) • tkinter (GUI)
Full Python code and GUI template available — estimated 40-60 hours development

## Slide 12
LABVIEW IMPLEMENTATION
NI USB-6001 CONNECTIONS
ANALOG INPUT (AI0): Pressure transducer 4-20mA Through 250Ω → 1-5V ANALOG OUTPUT (AO0): E/P regulator control (if using) 0-5V → 4-20mA converter DIGITAL OUTPUT (P0.0): SSR for solenoid PWM DIGITAL OUTPUT (P0.1-P0.3): Stepper PUL/DIR/ENA DIGITAL INPUT (P1.0): E-stop button
LABVIEW VI STRUCTURE
Main.vi ├── Initialize.vi (DAQ setup) ├── AcquirePressure.vi (AI read) ├── PID_Control.vi (built-in PID) ├── StepperControl.vi (DO pulses) ├── PWM_Control.vi (software PWM) ├── StateMachine.vi (modes) ├── SafetyCheck.vi (limits) └── DataLog.vi (TDMS write) Use Producer-Consumer pattern with queued message handler
LABVIEW ADVANTAGES
• Built-in PID toolkit with auto-tune capability • TDMS binary logging (faster than CSV, smaller files) • DAQmx drivers handle all timing/buffering • Professional front panel design tools • Real-time plotting and trend display • Easy to hand off to future students • UMich has campus-wide license — no cost • Industry standard in research labs
Check with MSE department — may already have NI DAQ hardware available

## Slide 13
IMPLEMENTATION TIMELINE
Target: Complete system operational within 4-6 weeks
Week 1
PROCUREMENT
Order all components; submit POs; verify specs
Week 2
MECHANICAL
Install valves & fittings; pressure test with N₂
Week 3
ELECTRICAL
Wire sensors & actuators; verify signals
Week 4
SOFTWARE
Develop control code; bench test with air
Week 5-6
COMMISSION
CO₂ tests; PID tuning; validate ramp rates
Critical Path: Motorized valve assembly and HP solenoid — order in Week 1

## Slide 14
COMPARISON SUMMARY
Factor
Commercial (Emerson)
DIY System
Cost
$12,000 - $15,000
$2,500 - $3,500
Timeline
3-4 months
4-6 weeks
Rate Accuracy
±2-3%
±5-10%
Data Logging
Limited (proprietary)
Full (custom format)
Customization
Limited
Unlimited
Support
Emerson tech support
Self + community
Learning Value
Minimal
Significant
Upgrade Path
Emerson ecosystem
Open platform
RECOMMENDATION: DIY system provides 80% of performance at 20% of cost

## Slide 15
SAFETY CONSIDERATIONS
HAZARDS
• Overpressure → vessel rupture • Rapid depressurization → dry ice, cold burns • Fitting failure → high-velocity jet • CO₂ accumulation → asphyxiation risk • Electrical → 24VDC, solenoid coils
MITIGATIONS
• Relief valve at 30 MPa (exists on vessel) • Software limits at 28 MPa (setpoint max) • Controlled vent rate via PWM solenoid • Fume hood operation during vent • Hardwired E-stop (power cutoff) • Torque all fittings to spec
REQUIRED SAFETY EQUIPMENT
• Mechanical relief valve (30 MPa set) — VERIFY EXISTING ON VESSEL • Rupture disc (32 MPa backup) — RECOMMENDED TO ADD • Check valve on booster output — PREVENTS BACKFLOW • Independent analog pressure gauge — BACKUP TO TRANSDUCER (already have) • E-stop button with power cutoff — ADD TO CONTROL PANEL • CO₂ monitor with 5000 ppm alarm — CHECK LAB EQUIPMENT
Fail-Safe Design: Power loss → booster stops, solenoid closes, vessel holds pressure

## Slide 16
NEXT STEPS
1
Confirm requirements with Dandan — rate accuracy tolerance?
2
Approve budget ($2,500 - $3,500) with professor
3
Order critical components this week (motorized valve, solenoid)
4
Select controller platform (Raspberry Pi vs LabVIEW)
5
Check MSE for available DAQ hardware
6
Begin mechanical installation in Week 2
Questions?

## Slide 17
APPENDIX: SUPPLIER CONTACTS
Component
Supplier
Contact / Notes
HP Solenoid
Clark Cooper (Magnatrol)
(973) 887-6400 | clarkcooper.com
E/P Regulator
SMC Corporation
(800) 762-7621 | smcusa.com
Pressure Transducer
Omega Engineering
(800) 826-6342 | omega.com
Swagelok Fittings
Michigan Fluid Power
(734) 213-8000 | Local pickup
NI DAQ
National Instruments
(888) 280-7645 | ni.com
Raspberry Pi
Digi-Key / Adafruit
digikey.com | adafruit.com
Stepper Motors
StepperOnline
omc-stepperonline.com
CO₂ Gas
Cryogenic Gases
(313) 835-5513 | Already using
Note: Get quotes before ordering — prices may vary. Check for UMich vendor agreements.