## Full BOM
Section | Item | Part Number | Description | Spec / Rating | Vendor | Qty | Unit Price | Line Total | Buy / Search Link
1. Motorized Needle Valve | Needle valve body | SS-1RS4 | Swagelok needle valve, regulating stem, 1/4" tube compression | 316SS, 5,000 PSI, Cv = 0.37 | idealvac.com / Swagelok distributor | 1 | 300 | 300 | Search this part
1. Motorized Needle Valve | Shaft coupler (NEMA 17 to valve stem) | 5mm to 1/4" flexible shaft coupler | Couples existing NEMA 17 stepper directly to valve stem | Aluminum flexible coupling, 5mm x 1/4" | Amazon | 1 | 10 | 10 | Search this part
2. Check Valve | Check valve | SS-CHS4-5 | Swagelok check valve, 1/4" tube fitting, 5 PSI cracking pressure | 316SS, 6,000 PSI (NOT SS-4C — that is only 3,000 PSI) | Amazon / Radwell.com | 1 | 100 | 100 | Search this part
3. Manual Ball Valves | Ball valve | SS-83KS4 | Swagelok 83 series ball valve, 1/4" tube fitting, PCTFE seats, CO2 compatible | 316SS, 6,000 PSI (NOT SS-43GS4 — that is only 3,000 PSI) | Radwell.com / eBay (surplus) | 3 | 407 | 1221 | Search this part
4. Stainless Tubing | SS tubing, 1/4" OD x 0.065" wall | SS-T4-S-065-20 | Swagelok seamless 316SS tubing, 20 ft minimum order | 316SS seamless, 1/4" OD x 0.065" wall, 20 ft | mcmaster.com | 1 | 120 | 120 | Search this part
5. Compression Fittings | Union (straight) | SS-400-6 | Swagelok union fitting, 1/4" tube OD | 316SS, 5,100 PSI rated | Amazon / Walmart | 4 | 15 | 60 | Search this part
5. Compression Fittings | 90 degree elbow | SS-400-9 | Swagelok 90 degree elbow fitting, 1/4" tube OD | 316SS, 5,100 PSI rated | Amazon / Walmart | 3 | 20 | 60 | Search this part
5. Compression Fittings | Tee | SS-400-3 | Swagelok tee fitting, 1/4" tube OD | 316SS, 5,100 PSI rated | Amazon / Walmart | 2 | 28 | 56 | Search this part
5. Compression Fittings | End cap | SS-400-C | Swagelok end cap fitting, 1/4" tube OD | 316SS, 5,100 PSI rated | Amazon | 3 | 10 | 30 | Search this part
6. Relief Valve | Relief valve | SS-RL3S4 | Swagelok high-pressure relief valve, 1/4" tube compression fitting ⚠ MUST specify set pressure ~4,350 PSI (30 MPa) when ordering ⚠ Previous BOM listed Parker 442F42 — WRONG (only rated 114 PSI max) | 316SS, 1/4" tube, rated to 6,000 PSI — matches existing Swagelok fittings Set between operating pressure (4,061 PSI) and vessel MAWP (4,999 PSI) Recommend set pressure: 4,350 PSI (30 MPa) | Swagelok authorized distributor | 1 | 250 | 250 | Search this part
6. Relief Valve | PTFE thread tape | 34P209 | PTFE sealing tape for NPT thread connections | High-pressure / high-temp rated | Grainger | 1 | 5 | 5 | Search this part
6. Relief Valve | Vent solenoid valve (automated depressurization) | Parker 4-Way or Asco 8290 HP | Normally-closed 24VDC solenoid valve, 1/4" tube or NPT, high-pressure rated Driven by RPi GPIO 18 (PWM) via relay — opens automatically during DEPRESSURIZE state ⚠ MUST be rated ≥ 6,000 PSI — standard solenoids (150-300 PSI) will fail catastrophically | 316SS body, NC (fail-safe closed), 24VDC coil, 1/4" process connection Candidates: Parker Series 34 HP, Asco 8290 HP series, or HiP solenoid valve ⚠ THIS PART IS MISSING FROM ORIGINAL BOM — confirm spec with distributor before ordering | Parker / Asco / High Pressure Equipment (HiP) | 1 | 500 | 500 | Search this part
7. Sensors | Pressure transducer | 5DEK9 | Ashcroft G2, 0-5000 PSI, 4-20 mA output, nylon housing, 316SS wetted parts, IP67 Grainger item 5DEK9 — Mfr model G17M0242F25000# ⚠ Do NOT order K4708 — that is the 1-5V DC output version, incompatible with 4-20mA wiring | 4-20 mA → 1-5 V via 250 ohm shunt resistor into ADS1115 A0 Nylon housing is fine for lab use; 316SS diaphragm is the CO2-wetted part | Grainger | 1 | 200 | 200 | Search this part
7. Sensors | Pressure gauge (mechanical, visual backup) | K4201 | Ashcroft 0-5000 PSI analog pressure gauge | Visual reference gauge, independent of electronics | Grainger | 1 | 100 | 100 | Search this part
8. Electronics & Control | Raspberry Pi 4 (4 GB) | RPi4-4GB | Single-board computer, runs the 10 Hz control loop in Python 3.11 | 4 GB RAM model recommended | Amazon / Adafruit | 1 | 55 | 55 | Search this part
8. Electronics & Control | ADS1115 16-bit I2C ADC module | ADS1115 | Reads pressure transducer (A0) and temperature sensor (A1) | 16-bit resolution, I2C interface | Adafruit / Amazon | 1 | 10 | 10 | Search this part
8. Electronics & Control | NEMA 17 stepper motor | NEMA17-200 | Actuates the motorized needle valve via shaft coupler | 200 steps/rev, standard NEMA 17 frame | Amazon | 1 | 15 | 15 | Search this part
8. Electronics & Control | A4988 stepper driver | A4988 | Drives NEMA 17 at 1/16 microstepping; GPIO 17 STEP, 27 DIR, 22 EN | 1/16 microstepping driver module | Pololu / Amazon | 1 | 8 | 8 | Search this part
8. Electronics & Control | 32 GB microSD card | microSD-32GB | Boot media for Raspberry Pi 4 | Class 10 / A1 rated for reliability | Amazon (SanDisk) | 1 | 10 | 10 | Search this part
8. Electronics & Control | Relay module, 24VDC | Relay-24VDC | Drives vent solenoid from GPIO 18 PWM signal | 24VDC coil, panel/DIN mountable | Amazon | 1 | 10 | 10 | Search this part
8. Electronics & Control | DIN rail relay, 24VDC coil | DIN-Relay-24VDC | DIN-rail mounted relay for panel-style wiring | 24VDC coil | Grainger | 1 | 20 | 20 | Search this part
8. Electronics & Control | DIN rail power supply, 24VDC 50W | 33NT20 | Dayton DIN rail PSU — powers relay, solenoid, and sensor circuits | 24VDC, 50W output | Grainger | 1 | 55 | 55 | Search this part
[7 empty] | SUBTOTAL (to buy) | 3195
9. Already Owned | Pressure vessel | PARR 2302HC | 316SS pressure vessel, S/N 4540-1803-78845A, 1 L volume | MAWP 5,000 PSI (34.47 MPa) @ 350°C — no purchase needed | N/A — already owned |  |  | $0.00 | N/A
9. Already Owned | Gas booster pump | HII 5G-TD-28/150-CO2 | Air-driven, twin-drive CO2 service booster, S/N 1205101 | Max outlet 25,000 PSIG (172 MPa) — no purchase needed | N/A — already owned |  |  | $0.00 | N/A
9. Already Owned | Temperature controller | INKBIRD | 60C setpoint temperature controller | ±0.01°C settled stability — no purchase needed | N/A — already owned |  |  | $0.00 | N/A
9. Already Owned | CO2 gas cylinder | UN1013 Bone Dry | Primary process gas — Metro Welding Supply, Detroit MI | No purchase needed (existing supply) | N/A — already owned |  |  | $0.00 | N/A
9. Already Owned | N2 gas cylinder | UN1066 Ultra High Purity | Alternate experiment gas — Metro Welding Supply, Detroit MI | No purchase needed (existing supply) | N/A — already owned |  |  | $0.00 | N/A
9. Already Owned | Ar gas cylinder | UN1006 Prepurified | Alternate experiment gas — Metro Welding Supply, Detroit MI | No purchase needed (existing supply) | N/A — already owned |  |  | $0.00 | N/A
9. Already Owned | Air gas cylinder (drive air) | UN1002 Dry Grade | Drives booster pump only — never enters vessel | No purchase needed (existing supply) | N/A — already owned |  |  | $0.00 | N/A
10. Optional: Windowed Vessel | Windowed pressure vessel (sapphire window) | EZE-Seal series | Autoclave Engineers (Parker) windowed reactor for flow visualization | Sapphire window, up to 60 MPa | Autoclave Engineers (Parker) | 1 | 5000 | 5000 | Search this part
10. Optional: Windowed Vessel | Windowed pressure vessel (alternate) | Sitec 100-600 mL | Sitec Reactor Technology windowed reactor, sapphire/borosilicate | Up to 100 MPa, 100-600 mL volume | Sitec Reactor Technology | 1 | 4500 | 4500 | Search this part
[7 empty] | TOTAL (required, excl. optional window vessel) | 3195

## Notes
scCO2 System — Complete BOM Notes
LINK VERIFICATION DISCLAIMER:
Automated link-checking was not possible in the environment this file was generated in
(all outbound web requests returned 403, including to sites like Wikipedia — a sandbox
network restriction, not a sign the URLs are wrong). Each link uses the vendor's standard,
publicly documented search-page URL format with the part number/description as the query
string — not a guessed deep product page. Please click a few to confirm before bulk-ordering.
Why these exact part numbers matter:
- SS-CHS4-5 (check valve) is rated 6,000 PSI. SS-4C looks similar but is only 3,000 PSI.
- SS-83KS4 (ball valve) is rated 6,000 PSI. SS-43GS4 looks similar but is only 3,000 PSI.
- All fittings are Swagelok SS-400 series, rated 5,100 PSI, matched to 1/4" tube OD.
- SS-1RS4 needle valve (5,000 PSI) gives ~23% margin above 4,061 PSI (28 MPa) operating pressure.
CORRECTIONS vs ORIGINAL BOM:
- Pressure transducer changed from K4708 to 5DEK9 (G17M0242F25000#). K4708 is the 1-5V DC
output variant — incompatible with the 4-20mA → 250Ω shunt wiring already designed.
5DEK9 is the correct 4-20mA version at roughly half the price (~$200 vs ~$400).
- Vent solenoid valve added (was missing from original BOM). Must be rated ≥6,000 PSI.
Standard 150-300 PSI solenoids WILL FAIL at 28 MPa. See Section 6 entry.
Section 9 (Already Owned) lists existing lab equipment for completeness — no purchase needed.
Section 10 (Optional) is a windowed vessel upgrade for visualizing scCO2 flow — not required
for the current closed-vessel system.
Recommended buying order:
1. Confirm thread/tube size compatibility with the existing PARR vessel fittings before ordering.
2. Order Swagelok parts from an authorized distributor when possible (warranty/traceability
matters on a 28 MPa system).
3. McMaster-Carr ships fastest for tubing; Grainger for relief valve, gauges, and PSU.
4. Electronics (Pi, ADC, stepper driver) are low-risk — any reputable seller is fine.